## Slide 1: Lensless Actinic Inspection of EUV Reticles with RESCAN
I. Mochi1, P. Helfenstein1, R. Rajendran1, I. Mohacsi1, Y. Ekinci1, S. Yoshitake2
1Paul Scherrer Institut. 2NuFlare Technology, Inc.
EUVL symposium 2016

## Slide 2: EUV reticle inspection for N7 and beyond
I. Mochi et al. Proc. Of SPIE Vol. 7636 76361A-1
Reticle SEM
AIT EUV image
DUV image
Wafer SEM
1 μm
0.5 μm
  Actinic inspection is the most accurate approach for the detection and characterization of defect on EUV reticles.
  No actinic patterned reticle inspection platform available for industrial use

## Slide 3: OUTLINE
Introduction to the RESCAN project
Coherent diffraction
Defect detection and localization
Experimental results
Low resolution defect map
Defect area image reconstruction
Rescan recent upgrades
A faster detector
Dedicated beamline
What’s next

## Slide 4: What is RESCAN?
  An actinic inspection platform for the localization and classification of defects on EUV reticles @ SLS synchrotron.
  Target specifications:
  Defect resolution: 10 nm
  Inspection throughput: 7h/reticle
  Based on coherent scattering methods

## Slide 5: How does RESCAN work?
Diffraction pattern
EUV reticle
Coherent
illumination
Detector
COHERENT SCATTERING

## Slide 6: A TWO-STEP PROCESS
THE DEFECTS ARE DETECTED AND LOCALIZED IN TWO STEPS
1. Fast mask scanning with variable-size probe
Diffraction pattern
analysis
2. Aerial image reconstruction of defect zone
CDI
reconstruction

## Slide 7: Defect identification
Reticle GDS
Calculated
Diffraction intensity
Recorded
Reference die
Low-resolution
defect map
Die to database comparison.
Die to die comparison.
High sensitivity.
Defect multiple detection.
Fast mask scan.
On the fly defect detection.

## Slide 8: Defect localization
Aerial image reconstruction with coherent diffraction imaging.
P-MA-06 Dr. P.Helfenstein Actinic EUV Mask Metrology with RESCAN using a monochromatic beam.
Iterative process.
Noise robust.
Magnitude and phase reconstruction.
0.5 μm

## Slide 9: Defect detection– experimental results
Page 9
Large added particle
25 nm hp grating
-40 x 2000 nm2
CD error
+40 x 2000 nm2
CD error

## Slide 10: Defect localization – experimental results
October 2015
February 2016
May 2016
July 2016
August 2016
June 2016
Artifact reduction.
Resolution enhancement.
Field of view increase.
Reconstruction algorithm optimization
RESCAN 1
RESCAN 2

## Slide 11: Defect localization – experimental results
Amplitude map
Amplitude modulation
100 nm CD lines with programmed and native defects

## Slide 12: Defect localization – experimental results
SEM image
Phase map
Phase contrast is also useful for defect identification.

## Slide 13: Looks promising but…
There are two main challenges for successful implementation:
1. Detector dynamic range
2. Full mask inspection time
Defect signal level
Detector max
Intensity typ. 3000 ph
Noise level
Low intensity – Signal buried in noise
High intensity – Peak saturation
How much time does it take to scan a 100×100 mm2 area with a 30 μm diameter probe?
We need at least 50% overlapping
100×100 mm2 / 0.015×0.015 mm2
44000000 images!
Requires ~1800 Hz frame rate

## Slide 14: RESCAN dedicated beamline at SLS
Wavelength: 4 -14 nm
Photon rate: 1011 -1012 Ph/s (Coherent flux on sample)
Monochromaticity: 2000
Designed and optimized for RESCAN

## Slide 15: A –very– fast detector
JUNGFRAU 0.2
JUNGFRAU 0.2 – Developed at PSI for hard X-Ray applications (FEL)
Dynamic range | >106 (20 bit)
Module size | 80mm X 40mm
Pixel size | 75 microns
No of pixels | 512 × 1024
Acquisition Rate | 2000 Hz
Pixel Bleeding | No in both directions
P-MA-03 Dr. R.Rajendran A high frame-rate pixel detector with extended dynamic range for EUV imaging and mask metrology.

## Slide 16: Defect inspection at 500 Hz
100 nm hp grating
10 nm
CD error
X position [μm]
Y position [μm]
A.U.
Recent data.

## Slide 17
Multi-module detector
Up to six independent modules. No decrease in acquisition speed.
EUV optimized. Quantum efficiency increases from 3% to 50%.
Jungfrau detector upgrade

## Slide 18
Wavelength | 13.5 nm
Flux | >100 mW
Brilliance | > 106 W/(mm2⋅strd)
Beam energy/beam current | 430 MeV/150 mA
Pulse structure | ~50 ps every 2 ns
Injection system | Top-up mode
Footprint | 5m × 12m
Compact EUV Metrology Source „COSAMI“

## Slide 19: Conclusion
The RESCAN platform can be a viable approach to fill the gap in actinic reticle defect inspection.
The key technology is available today.
Additional component development and integration is in progress:
Fast detector
Improved algorithms for inspection and image reconstruction
Compact EUV source

## Slide 20: Wir schaffen Wissen – heute für morgen We create knowledge – today for tomorrow
Aknowledgements
M. Vockenhuber
M. Kropf
J. Gabadinho
A. Mozzanica
S. Redford
B. Schmitt