## Slide 1: Actinic inspection of programmed defects on arbitrary patterns
Iacopo Mochi, Patrick Helfenstein, Rajendran Rajeev, Sara Fernandez, Dimitrios Kazazis, Shusuke Yoshitake and Yasin Ekinci

## Slide 2: Outline
RESCAN – A platform for actinic pattern inspection
Description.
Working principle.
Non periodic patterns
Programmed defect sample preparation.
Die to die inspection results.
Die to database inspection results.
Undergoing hardware and software upgrades
High dynamic range sensor.
Improving the resolution.
Increasing the inspection speed.

## Slide 3: RESCAN
EUV reticle
Coherent EUV
illumination
Detector
Actinic pattern inspection platform based on Coherent Diffraction Imaging (CDI)
Synchrotron-based
Tunable wavelength
  EUV optimized
λ/∆ λ ≈ 1500
Max sample size: 2×2 cm2
Inspection area: 200x200 μm2
Namax: 0.24
Resolution limit: 34 nm
Compact vacuum chamber
XIL-II beamline SLS

## Slide 4: Current results – defects on periodic patterns
Iacopo Mochi et al. (2017). RESCAN: an actinic lensless microscope for defect inspection of EUV reticles. Journal of Micro/Nanolithography, MEMS, and MOEMS, 16(4), 041003–041003.
Defect inspection on periodic patterns based on coherent scattering
SNR: 6.9
SNR: 7.1
Non periodic structures are more challenging for a simple approach based on coherent scattering.
Notes: Therefore we use CDI… and specifically ptychography

## Slide 5: How does ptychography work?
The phase of the far field diffraction pattern is not recorded.
Ptychography iteratively searches for a common solution to the inverse problem in the overlapping area

## Slide 6: Programmed defect in a random pattern
Variable size extrusion defects
Pattern CD: 200 nm (on mask)

## Slide 7: Programmed defect sample
200 nm
201.8 nm
151.3 nm
100.9 nm
50.5 nm
Silicon substrate.
MoSi multilayer: 40 bi-layers.
HSQ absorber.
Thickness: 145 nm.
Defect size on mask
200×200 nm extrusion.
200×150 nm extrusion.
200×100 nm extrusion.
200×50 nm extrusion.
Dark-field sample with programmed extrusion defects.

## Slide 8: Pattern comparison for accurate defect localization
μm
Reference
μm
200-nm extrusion
Defect map
Reconstructed area: 20×20 μm2.
Average iteration time: 80 ms.
Defect SNR: 15
Die to die defect inspection:
comparing a defect free area of the sample to a defective one.

## Slide 9: Pattern comparison for accurate defect localization
μm
200-nm extrusion
Defect map
μm
Reference
Defect phase map
Strong phase contrast
Binary phase
rad
rad

## Slide 10: A practical approach: die to database inspection
μm
μm
Comparison between reconstructed image and mask design.
200×200 nm2 extrusion defect.
Apply spatial frequency filter to reticle design.
Calculate the best fitting affine transformation matrix.
Align and compare the images
Mask layout
Ideal mask image
Reconstructed amplitude image
Defect map

## Slide 11: Close to the resolution limit
200 nm – SNR 7.3
150 nm – SNR 6.2
100 nm – SNR 5.2
50 nm – SNR 6.5
μm
μm
μm
μm
Defects down to 50 nm (on mask) are detected with a significant SNR.
The current resolution limit of RESCAN is 34 nm.
Notes: To help the defect visualization we can threshold and binarize the images.

## Slide 12: Planned upgrades and roadmap
JUNGFRAU and COSAMI: Detector and source technologies for high-throughput EUV mask inspection, Rajeev Rajendran et al. [10451-49]
From EUV CCD to two-modules EUV-optimized JUNGFRAU detector. (10/2017)
Large dynamic range
Single acquisition and short exposure time (0.01 s).
High NA
Resolution limit down to 20 nm.
EUV CCD
Pixel size: 13.5 μm.
Acquisition frame rate: 1 Hz.
Dynamic range: 16 bit
NA: 0.24
Pixel number: 2048×2048
JUNGFRAU
Pixel number: 1024×512 (1 module)
Pixel size: 75μm.
Acquisition frame rate: 2000 Hz.
Dynamic range: 20 bit.
NA: 0.4

## Slide 13: Planned upgrades and roadmap
GPU integration for ptychography reconstruction.
Aerial image simulation for accurate die to database inspection.
Improved reconstruction code. (02/2018)
COmpact Source for Actinic Mask Inspection
JUNGFRAU and COSAMI: Detector and source technologies for high-throughput EUV mask inspection, Rajeev Rajendran et al. [10451-49]
Wavelength | 13.5 nm
Flux | >100 mW
Brilliance | > 106 W/mm2⋅strd
Beam Energy | 430 MeV
Beam Current | 150 mA
Pulse Structure | ~50 ps every 2 ns
Injection mode | Top-up mode
Foot print | 5m × 12m

## Slide 14: Summary
RESCAN can effectively map programmed defects on periodic and non periodic patterns.
Die-to-die and die-to-database inspection modes are available.
We are upgrading RESCAN with a high NA detector that will enable defect inspection down to 20 nm.
We demonstrated defect inspection down to 50 nm (on mask) with a high SNR.
We are upgrading the software to improve the reconstruction speed and stability.
10/2017
02/2018
RESCAN latest results and upgrades

## Slide 15: Wir schaffen Wissen – heute für morgen
Acknowledgements
Michaela Vockenhuber
Marcus Kropf
Luca Bühlmann
Zuhal Tasdemir